## DAQ Signals All
Subsystem | Name | Range | Output | Pins | Connector | Card | Rate (Hz) | Current PXI Slot | Description
PAS | Speaker Out | 0-10 V | AO0 | 55,21 | 0 | 7842R | 80k | 8 | Runs a speaker signal from 0 to 10 V at the maximum range.
| LaserOut 0 | -10-10 V | AO1 | 54,20 | 0 | 7842R | 40k | 8 | Laser drive signal for cell 0
| LaserOut 1 | -10-10 V | AO2 | 53,19 | 0 | 7842R | 40k | 8 | Laser drive signal for cell 1
| LaserOut 2 | -10-10 V | AO3 | 52,18 | 0 | 7842R | 40k | 8 | Laser drive signal for cell 2
| LaserOut 3 | -10-10 V | AO4 | 51,17 | 0 | 7842R | 40k | 8 | Laser drive signal for cell 3
| LaserOut 4 | -10-10 V | AO5 | 50,16 | 0 | 7842R | 40k | 8 | Laser drive signal for cell 4
| Speaker Enable | TTL | DIO10 | 11,9 | 0 | 7842R | N/A | 8 | Switch between speaker and laser output. High == speaker.
| iSpk Enable 0 | TTL | DIO0 | 36,8 | 0 | 7842R | N/A | 8 | Individual cell 0 enable signal for the speaker. High indicates that speaker is not powered.
| iSpk Enable 1 | TTL | DIO1 | 37,8 | 0 | 7842R | N/A | 8 | Individual cell 1 enable signal for the speaker. High indicates that speaker is not powered.
| iSpk Enable 2 | TTL | DIO2 | 38,8 | 0 | 7842R | N/A | 8 | Individual cell 2 enable signal for the speaker. High indicates that speaker is not powered.
| iSpk Enable 3 | TTL | DIO3 | 39,8 | 0 | 7842R | N/A | 8 | Individual cell 3 enable signal for the speaker. High indicates that speaker is not powered.
| iSpk Enable 4 | TTL | DIO4 | 40,8 | 0 | 7842R | N/A | 8 | Individual cell 4 enable signal for the speaker. High indicates that speaker is not powered.
| Laser Enable 0 | TTL | DIO5 | 41,8 | 0 | 7842R | N/A | 8 | Laser enable signal for cell 0. High == laser powered.
| Laser Enable 1 | TTL | DIO6 | 42,8 | 0 | 7842R | N/A | 8 | Laser enable signal for cell 1. High == laser powered.
| Laser Enable 2 | TTL | DIO7 | 43,8 | 0 | 7842R | N/A | 8 | Laser enable signal for cell 2. High == laser powered.
| Laser Enable 3 | TTL | DIO8 | 10,8 | 0 | 7842R | N/A | 8 | Laser enable signal for cell 3. High == laser powered.
| Laser Enable 4 | TTL | DIO9 | 44,8 | 0 | 7842R | N/A | 8 | Laser enable signal for cell 4. High == laser powered.
| Microphone 0 |  | AI0 | 68,34 | 0 | 7842R | 40k | 8 | Microphone signal for cell 0
| Microphone 1 |  | AI1 | 66,32 | 0 | 7842R | 40k | 8 | Microphone signal for cell 1
| Microphone 2 |  | AI2 | 65,31 | 0 | 7842R | 40k | 8 | Microphone signal for cell 2
| Microphone 3 |  | AI3 | 63,29 | 0 | 7842R | 40k | 8 | Microphone signal for cell 3
| Microphone 4 |  | AI4 | 62,28 | 0 | 7842R | 40k | 8 | Microphone signal for cell 4
| Photodiode 0 |  | AI0 | 0,8 | 0 | 6225 | 40k | 6 | Photodiode signal for cell 0
| Photodiode 1 |  | AI1 | 1,9 | 0 | 6225 | 40k | 6 | Photodiode signal for cell 1
| Photodiode 2 |  | AI2 | 2,10 | 0 | 6225 | 40k | 6 | Photodiode signal for cell 2
| Photodiode 3 |  | AI3 | 3,11 | 0 | 6225 | 40k | 6 | Photodiode signal for cell 3
| Photodiode 4 |  | AI4 | 4,12 | 0 | 6225 | 40k | 6 | Photodiode signal for cell 4
| Laser RMS 0 |  | AI5 | 5,13 | 0 | 6225 | 40k | 6 | Laser RMS signal for cell 0
| Laser RMS 1 |  | AI6 | 6,14 | 0 | 6225 | 40k | 6 | Laser RMS signal for cell 1
| Laser RMS 2 |  | AI16 | 16,24 | 1 | 6225 | 40k | 6 | Laser RMS signal for cell 2
| Laser RMS 3 |  | AI17 | 17,25 | 1 | 6225 | 40k | 6 | Laser RMS signal for cell 3
| Laser RMS 4 |  | AI18 | 18,26 | 1 | 6225 | 40k | 6 | Laser RMS signal for cell 4
| Tcell_00 |  | AI19 | 19,27 | 1 | 6225 | 40k | 6 | Cell 0, thermistor 0
| Tcell_01 |  | AI20 | 20,28 | 1 | 6225 | 40k | 6 | Cell 0, thermistor 1
| Tcell_10 |  | AI21 | 21,29 | 1 | 6225 | 40k | 6 | Cell 1, thermistor 0
| Tcell_11 |  | AI22 | 22,30 | 1 | 6225 | 40k | 6 | Cell 1, thermistor 1
| Tcell_20 |  | AI23 | 23,31 | 1 | 6225 | 40k | 6 | Cell 2, thermistor 0
| Tcell_21 |  | AI32 | 32,40 | 1 | 6225 | 40k | 6 | Cell 2, thermistor 1
| Tcell_30 |  | AI33 | 33,41 | 1 | 6225 | 40k | 6 | Cell 3, thermistor 0
| Tcell_31 |  | AI34 | 34,42 | 1 | 6225 | 40k | 6 | Cell 3, thermistor 1
| Tcell_40 |  | AI35 | 35,43 | 1 | 6225 | 40k | 6 | Cell 4, thermistor 0
| Tcell_41 |  | AI36 | 36,44 | 1 | 6225 | 40k | 6 | Cell 4, thermistor 1
| Denuded Bypass |  | P0.0 | 52,18 | 0 | 6225 | On Demand | 6 | These valves control the flow through the denuded channel of the PAS - whether the flow will bypass the denuder or go through it.
|  |  | P0.1 | 17,18 | 0 |  |  | 6
General | Filter CMD |  | P0.0 |  | 0 | 6225 | On Demand | 7 | These valves control whether the flow will be sample or through the filter.
|  |  | P0.1 |  | 0 |  |  | 7
| Filter Rtn |  | P0.5 |  | 0 |  |  | 7 | These digital IO are feedback for current filter position.
|  |  | P0.6 |  | 0 |  |  | 7
| Cabin CMD |  | P0.2 |  | 0 | 6225 | On Demand | 7 | These valves control where the main flow comes from - through the pressure controller and cabin inlet or through the aerosol inlet.
|  |  | P0.3 |  | 0 |  |  | 7
| Cabin Rtn |  | P1.0 |  | 0 |  |  | 7 | These DIO are feedback for the current cabin valve position.
|  |  | P1.1 |  | 0 |  |  | 7
| O3 |  | P0.4 |  | 0 | 6225 | On Demand | 7 | O3 addition valve
| Inlet LED |  | P1.4 | 41,7 | 0 | 6225 | On Demand | 6
| Filter LED |  | P2.1 | 3,4 | 0 | 6225 | On Demand | 6
| Denuder LED |  | P2.2 | 45,12 | 0 | 6225 | On Demand | 6
| O3 LED |  | P2.3 | 46,12 | 0 | 6225 | On Demand | 6
| Spk LED |  | P2.4 | 2,4 | 0 | 6225 | On Demand | 6
| 405 LED |  | P2.5 | 40,7 | 0 | 6225 | On Demand | 6
| 515 LED |  | P2.6 | 1,35 | 0 | 6225 | On Demand | 6
| 660 LED |  | P2.7 | 39,4 | 0 | 6225 | On Demand | 6
| Error |  | P0.2 | 49,15 | 0 | 6225 | On Demand | 6
| Heartbeat |  | P0.3 | 47,13 | 0 | 6225 | On Demand | 6
| Sample Pump Enable |  | P0.4 | 19,53 | 0 | 6225 | On Demand | 6
| O3 Generator Enable |  | P0.5 | 51,18 | 0 | 6225 | On Demand | 6
| Denuder Enable |  | P0.6 | 16,50 | 0 | 6225 | On Demand | 6
| Laser PS Enable |  | P0.7 | 48,15 | 0 | 6225 | On Demand | 6
| TEC PS Enable |  | P1.0 | 11,44 | 0 | 6225 | On Demand | 6
| Sample Pump Block |  | P1.4 | 10,44 | 0 | 6225 | On Demand | 7 | Input from sample pump indicating blockage
| Laser Interlock |  | P1.2 | 43,9 | 0 | 6225 | On Demand | 7 | Input from laser enclosure interlock
| Blocked Impactor |  | P1.3 | 42,8 | 0 | 6225 | On Demand | 7 | Signal from pressure differential sensor to indicate blocked impactor
| Ozone DO |  | ctr0 | 37,4 | 0 | 6225 | On Demand | 6 | Output for PWM of ozone generator
CRDS | CRDS Blue AI |  | AI0 | 68,34 | 0 | 6124 |  | 3 | Analog input for CRDS blue, dry PMT
|  |  | AI1 | 33,66 | 0 | 6124 |  | 3 | Analog input for CRDS blue, dry , filtered PMT
|  |  | AI2 | 65,31 | 0 | 6124 |  | 3 | Analog input for CRDS blue, high RH PMT
|  |  | AI3 | 30,63 | 0 | 6124 |  | 3 | Analog input for CRDS blue, med RH PMT
| CRDS Red AI |  | AI0 | 68,34 | 0 | 6132 |  | 5 | Analog input for CRDS red PMT
| CRD Blue Laser Output |  | Ctr0 | 2,36 | 0 | 6225 |  | 8 | Modulation for blue diode
| CRD Red Laser Output |  | Ctr1 | 40,7 | 0 | 6225 |  | 8 | Modulation for red diode
| PMT Gain |  | AO5 | 60,27 | N/A | 6704 |  | 6
|  |  | AO6 | 25,59 | N/A | 6704 |  | 6
|  |  | AO7 | 57,24 | N/A | 6704 |  | 6
|  |  | AO8 | 22,55 | N/A | 6704 |  | 6
|  |  | AO9 | 54,20 | N/A | 6704 |  | 6
| Red Enable |  | P0.0 |  |  | 6225 |  | 7 | Laser enable signal for red laser on CRDS
| Blue Enable |  | P0.1 |  | 0 | 6225 |  | 7 | Laser enable signal for blue laser 0 on CRDS
|  |  | P0.2 |  | 0 | 6225 |  | 7 | Laser enable signal for blue laser 1 on CRDS
| Red Laser Gain |  | AO0 |  |  | 6704
| Blue0 Gain |  | AO1 |  |  | 6704
| Blue1 Gain |  | AO2 |  |  | 6704

## Digital Signals All
Signal | Type | Port | Address | Device | Description
PAS TECs | 485 |  | 01 | TEC-1090 | Cell 1 TEC
|  |  | 02 | TEC-1090 | Cell 2 TEC
|  |  | 03 | TEC-1090 | Cell 3 TEC
|  |  | 04 | TEC-1090 | Cell 4 TEC
|  |  | 05 | TEC-1090 | Cell 5 TEC
CRD TECs |  |  | 06 | TEC-1090 | Med RH TEC
|  |  | 07 | TEC-1090 | High RH TEC
Mass Flow Controllers | 232 | COM0 | A | MC-5SLPM | CRDS Cell 1, PAS Cell 5
|  |  | B | MC-5SLPM | CRDS Cell 4, PAS Cell 4
|  |  | C | MC-5SLPM | PAS Cell 1
|  |  | D | MC-5SLPM | PAS Cell 2
|  |  | E | MC-5SLPM | PAS Cell 3
|  |  | F | MC-5SLPM | CLAP
|  |  | G | MC-5SLPM | CRDS Cell 2
|  |  | H | MC-5SLPM | CRDS Cell 5
|  |  | J | MC-1SLPM | CRDS Mirror Purge
|  |  | K | MC-1SLPM | O3 Bypass
PAS Cooling | 485 |  |  | TC-48-20 | TE-Technology TEC cooling PAS baseplate
Pressure Transducers | 485 |  | 01 | PPT0015AXN5VB | Drier PPT
|  |  | 02 | PPT0015AXN5VB | CRD Humidified
|  |  | 03 | PPT0015AXN5VB | Dry pressure
Hygrometers | 485 |  | 0 | HMP60 | Wet side of humidifier
|  |  | 1 | HMP60 | Dry side of humidifier
|  |  | 2 | HMP60 | Outlet of CRDS Cell 2
|  |  | 3 | HMP60 | Outlet of CRDS Cell 5
|  |  | 4 | HMP60 | PAS Cell 1
|  |  | 5 | HMP60 | PAS Cell 2
|  |  | 6 | HMP60 | PAS Cell 3
|  |  | 7 | HMP60 | PAS Cell 4
|  |  | 8 | HMP60 | PAS Cell 5

## PAS
Name | Range | Output | Pins | Connector | Card | Rate (Hz) | Current PXI Slot | Tested? | Description
Speaker Out | 0-10 V | AO0 | 55,21 | 0 | 7842R | 80k | 8 | Yes - Test FPGA | Runs a speaker signal from 0 to 10 V at the maximum range.
LaserOut 0 | -10-10 V | AO1 | 54,20 | 0 | 7842R | 40k | 8 | Yes - Test FPGA | Laser drive signal for cell 0
LaserOut 1 | -10-10 V | AO2 | 53,19 | 0 | 7842R | 40k | 8 | Yes - Test FPGA | Laser drive signal for cell 1
LaserOut 2 | -10-10 V | AO3 | 52,18 | 0 | 7842R | 40k | 8 | Yes - Test FPGA | Laser drive signal for cell 2
LaserOut 3 | -10-10 V | AO4 | 51,17 | 0 | 7842R | 40k | 8 | Yes - Test FPGA | Laser drive signal for cell 3
LaserOut 4 | -10-10 V | AO5 | 50,16 | 0 | 7842R | 40k | 8 | Yes - Test FPGA | Laser drive signal for cell 4
Speaker Enable | TTL | DIO10 | 11,9 | 0 | 7842R | N/A | 8 | Yes - Test FPGA | Switch between speaker and laser output. High == speaker.
iSpk Enable 0 | TTL | DIO0 | 36,8 | 0 | 7842R | N/A | 8 | Yes - Test FPGA | Individual cell 0 enable signal for the speaker. High indicates that speaker is not powered.
iSpk Enable 1 | TTL | DIO1 | 37,8 | 0 | 7842R | N/A | 8 | Yes - Test FPGA | Individual cell 1 enable signal for the speaker. High indicates that speaker is not powered.
iSpk Enable 2 | TTL | DIO2 | 38,8 | 0 | 7842R | N/A | 8 | Yes - Test FPGA | Individual cell 2 enable signal for the speaker. High indicates that speaker is not powered.
iSpk Enable 3 | TTL | DIO3 | 39,8 | 0 | 7842R | N/A | 8 | Yes - Test FPGA | Individual cell 3 enable signal for the speaker. High indicates that speaker is not powered.
iSpk Enable 4 | TTL | DIO4 | 40,8 | 0 | 7842R | N/A | 8 | Yes - Test FPGA | Individual cell 4 enable signal for the speaker. High indicates that speaker is not powered.
Laser Enable 0 | TTL | DIO5 | 41,8 | 0 | 7842R | N/A | 8 | Yes - Test FPGA | Laser enable signal for cell 0. High == laser powered.
Laser Enable 1 | TTL | DIO6 | 42,8 | 0 | 7842R | N/A | 8 | Yes - Test FPGA | Laser enable signal for cell 1. High == laser powered.
Laser Enable 2 | TTL | DIO7 | 43,8 | 0 | 7842R | N/A | 8 | Yes - Test FPGA | Laser enable signal for cell 2. High == laser powered.
Laser Enable 3 | TTL | DIO8 | 10,8 | 0 | 7842R | N/A | 8 | Yes - Test FPGA | Laser enable signal for cell 3. High == laser powered.
Laser Enable 4 | TTL | DIO9 | 44,8 | 0 | 7842R | N/A | 8 | Yes - Test FPGA | Laser enable signal for cell 4. High == laser powered.
Microphone 0 |  | AI0 | 68,34 | 0 | 7842R | 40k | 8 | Yes - Test FPGA | Microphone signal for cell 0
Microphone 1 |  | AI1 | 66,32 | 0 | 7842R | 40k | 8 | Yes - Test FPGA | Microphone signal for cell 1
Microphone 2 |  | AI2 | 65,31 | 0 | 7842R | 40k | 8 | Yes - Test FPGA | Microphone signal for cell 2
Microphone 3 |  | AI3 | 63,29 | 0 | 7842R | 40k | 8 | Yes - Test FPGA | Microphone signal for cell 3
Microphone 4 |  | AI4 | 62,28 | 0 | 7842R | 40k | 8 | Yes - Test FPGA | Microphone signal for cell 4
Photodiode 0 |  | AI0 | 0,8 | 0 | 6225 | 40k | 6 |  | Photodiode signal for cell 0
Photodiode 1 |  | AI1 | 1,9 | 0 | 6225 | 40k | 6 |  | Photodiode signal for cell 1
Photodiode 2 |  | AI2 | 2,10 | 0 | 6225 | 40k | 6 |  | Photodiode signal for cell 2
Photodiode 3 |  | AI3 | 3,11 | 0 | 6225 | 40k | 6 |  | Photodiode signal for cell 3
Photodiode 4 |  | AI4 | 4,12 | 0 | 6225 | 40k | 6 |  | Photodiode signal for cell 4
Laser RMS 0 |  | AI5 | 5,13 | 0 | 6225 | 40k | 6 |  | Laser RMS signal for cell 0
Laser RMS 1 |  | AI6 | 6,14 | 0 | 6225 | 40k | 6 |  | Laser RMS signal for cell 1
Laser RMS 2 |  | AI16 | 16,24 | 1 | 6225 | 40k | 6 |  | Laser RMS signal for cell 2
Laser RMS 3 |  | AI17 | 17,25 | 1 | 6225 | 40k | 6 |  | Laser RMS signal for cell 3
Laser RMS 4 |  | AI18 | 18,26 | 1 | 6225 | 40k | 6 |  | Laser RMS signal for cell 4
Tcell_00 |  | AI19 | 19,27 | 1 | 6225 | 40k | 6 |  | Cell 0, thermistor 0
Tcell_01 |  | AI20 | 20,28 | 1 | 6225 | 40k | 6 |  | Cell 0, thermistor 1
Tcell_10 |  | AI21 | 21,29 | 1 | 6225 | 40k | 6 |  | Cell 1, thermistor 0
Tcell_11 |  | AI22 | 22,30 | 1 | 6225 | 40k | 6 |  | Cell 1, thermistor 1
Tcell_20 |  | AI23 | 23,31 | 1 | 6225 | 40k | 6 |  | Cell 2, thermistor 0
Tcell_21 |  | AI32 | 32,40 | 1 | 6225 | 40k | 6 |  | Cell 2, thermistor 1
Tcell_30 |  | AI33 | 33,41 | 1 | 6225 | 40k | 6 |  | Cell 3, thermistor 0
Tcell_31 |  | AI34 | 34,42 | 1 | 6225 | 40k | 6 |  | Cell 3, thermistor 1
Tcell_40 |  | AI35 | 35,43 | 1 | 6225 | 40k | 6 |  | Cell 4, thermistor 0
Tcell_41 |  | AI36 | 36,44 | 1 | 6225 | 40k | 6 |  | Cell 4, thermistor 1
TEC 0 |  | VCH0 | 34,68 | N/A | 6704 | On Demand | 6 |  | Temperature control for cell 0
TEC 1 |  | VCH1 | 66,33 | N/A | 6704 | On Demand | 6 |  | Temperature control for cell 1
TEC 2 |  | VCH2 | 31,65 | N/A | 6704 | On Demand | 6 |  | Temperature control for cell 2
TEC 3 |  | VCH3 | 63,30 | N/A | 6704 | On Demand | 6 |  | Temperature control for cell 3
TEC 4 |  | VCH4 | 28,62 | N/A | 6704 | On Demand | 6 |  | Temperature control for cell 4
TEC Enable 0 |  | DI0 | 2,36 | N/A | 6704 | On Demand | 6 |  | Temperature controller enable for cell 0. High is enabled.
TEC Enable 1 |  | DI1 | 3,36 | N/A | 6705 | On Demand | 6 |  | Temperature controller enable for cell 1. High is enabled.
TEC Enable 2 |  | DI2 | 4,36 | N/A | 6706 | On Demand | 6 |  | Temperature controller enable for cell 2. High is enabled.
TEC Enable 3 |  | DI3 | 5,36 | N/A | 6707 | On Demand | 6 |  | Temperature controller enable for cell 3. High is enabled.
TEC Enable 4 |  | DI4 | 6,36 | N/A | 6708 | On Demand | 6 |  | Temperature controller enable for cell 4. High is enabled.

## General
Name | Range | Output | Pins | Connector | Card | Current PXI Slot | Tested? | Description
Filter |  | P0.0 |  |  | 6225 | 7
|  | P0.1 |  |  | 6225 | 7
Cabin |  | P0.2 |  |  | 6225 | 7
|  | P0.3 |  |  | 6225 | 7
O3 |  | P0.4 |  |  | 6225 | 7

## CRDS
Name | Range | Output | Pins | Connector | Card | Current PXI Slot | Tested? | Description
Blue, Dry PMT In |  | AI0 | 68,34 | 0 | PXIe-6124 | 3
Blue, Dry Filtered PMT |  | AI1 | 33,66 | 0 | PXIe-6124 | 3
Blue, High RH PMT |  | AI2 | 65,31 | 0 | PXIe-6124 | 3
Blue, Med RH PMT |  | AI3 | 30,63 | 0 | PXIe-6124 | 3
Red, Dry PMT |  | AI0 | 68,34 | 0 | PXI-6132 | 5
Blue Laser Output |  | Ctr0 | 2,36 | 0 | 6225 | 8 |  | Modulation for blue diode
Red Laser Output |  | Ctr1 | 40,7 | 0 | 6225 | 8 |  | Modulation for red diode
PMT Gain 0 |  | AO5 | 60,27 | N/A | 6704 | 6
PMT Gain 1 |  | AO6 | 25,59 | N/A | 6704 | 6
PMT Gain 2 |  | AO7 | 57,24 | N/A | 6704 | 6
PMT Gain 3 |  | AO8 | 22,55 | N/A | 6704 | 6
PMT Gain 4 |  | AO9 | 54,20 | N/A | 6704 | 6
Red, Enable |  | P0.0 |  |  | 6225 | 8
Blue 0 Enable |  | P0.1 |  | 0 | 6225 | 8
Blue 1 Enable |  | P0.2 |  | 0 | 6225 | 8
Denuded Bypass |  | P0.3 |  | 0 | 6225 | 8
|  | P0.4 |  | 0 | 6225 | 8

## Commands
Group | Signal | Units | Range | Format | Description
Flow Control | Qsp,a | lpm | 0-10 | Numeric | CRDS cell 1, PAS cell 5 flow control
| Qsp,b | lpm | 0-10 | Numeric | CRDS cell 4, PAS cell 4 flow control
| Qsp,c | lpm | 0-10 | Numeric | PAS cell 1 flow control
| Qsp,d | lpm | 0-10 | Numeric | PAS cell 2 flow control
| Qsp,e | lpm | 0-10 | Numeric | PAS cell 3 flow control
| Qsp,f | lpm | 0-10 | Numeric | CLAP flow control
| Qsp,g | lpm | 0-10 | Numeric | CRDS cell 2 flow control
| Qsp,h | lpm | 0-10 | Numeric | CRDS cell 5 flow control
| Qsp,j | lpm | 0-10 | Numeric | Mirror purge flow
| Qsp,k | lpm | 0-10 | Numeric | Ozone generator flow control
| Qsp,m | mb | 0-2000 | Numeric | Pressure controller flow control
CRDS | fblue | Hz | 0-3000 | Numeric | CRDS blue laser frequency
| DCblue | % | 0-100 | Numeric | CRDS blue laser duty cycle
| fred | HZ | 0-3000 | Numeric | CRDS red laser frequency
| DCred | % | 0-100 | Numeric | CRDS red laser duty cycle
| Av, 0 | V | 0-10 | Numeric | Gain for PMT on cell 0
| Av, 1 | V | 0-10 | Numeric | Gain for PMT on cell 1
| Av, 2 | V | 0-10 | Numeric | Gain for PMT on cell 2
| Av, 3 | V | 0-10 | Numeric | Gain for PMT on cell 3
| Av, 4 | V | 0-10 | Numeric | Gain for PMT on cell 4
| fAI,blue | MHz | 1-4 | Numeric | Acquisition rate for blue laser
| fAI,red | MHz | 1-2.5 | Numeric | Acquisition rate for red laser
| RH2 | % | 0-100 | Numeric | Humidity setpoint for CRDS cell 2
| RH5 | % | 0-100 | Numeric | Humidity setpoint for CRDS cell 5
| P2 |  | 0 - 100 | Numeric | Proportional gain for cell 2 controller
| I2 |  | 0 - 100 | Numeric | Integral gain for cell 2 TEC
| D2 |  | 0 - 100 | Numeric | Derivative gain for cell 2 TEC
| P5 |  | 0 - 100 | Numeric | Proportional gain for cell 5 TEC
| I5 |  | 0 - 100 | Numeric | Integral gain for cell 5 TEC
| D5 |  | 0 - 100 | Numeric | Derivative gain for cell 5 TEC
| Denuder Valve |  | BYPASS/DENUDED | Boolean | Path control for denuded channels
PAS | fres,0 | Hz | 0-5000 | Numeric | PAS laser modulation frequency for cell 0
| Mod0 | N/A | Sine/Square | Enumeration | PAS modulation method for cell 0
| Vrange, 0 | V | 0-20 | Numeric | PAS modulation laser range for cell 0
| Voffset,0 | V |  | Numeric | Pas modulation offset for cell 0
| fres,1 | Hz | 0-5000 | Numeric | PAS laser modulation frequency for cell 1
| Mod1 | N/A | Sine/Square | Enumeration | PAS modulation method for cell 1
| Vrange, 1 | V | 0-20 | Numeric | PAS modulation laser range for cell 1
| Voffset,1 | V |  | Numeric | Pas modulation offset for cell 1
| fres,2 | Hz | 0-5000 | Numeric | PAS laser modulation frequency for cell 2
| Mod2 | N/A | Sine/Square | Enumeration | PAS modulation method for cell 2
| Vrange, 2 | V | 0-20 | Numeric | PAS modulation laser range for cell 2
| Voffset,2 | V |  | Numeric | Pas modulation offset for cell 2
| fres,3 | Hz | 0-5000 | Numeric | PAS laser modulation frequency for cell 3
| Mod3 | N/A | Sine/Square | Enumeration | PAS modulation method for cell 3
| Vrange, 3 | V | 0-20 | Numeric | PAS modulation laser range for cell 3
| Voffset,3 | V |  | Numeric | Pas modulation offset for cell 3
| fres,4 | Hz | 0-5000 | Numeric | PAS laser modulation frequency for cell 4
| Mod4 | N/A | Sine/Square | Enumeration | PAS modulation method for cell 4
| Vrange, 4 | V | 0-20 | Numeric | PAS modulation laser range for cell 4
| Voffset,4 | V |  | Numeric | Pas modulation offset for cell 4
| Speaker Switch |  | Laser/Speaker | Enumeration | Switch for changing between laser and speaker
| Speaker Enable 0 |  | OFF/ON | Boolean | Speaker enable signal for PAS cell 0
| Speaker Enable 1 |  | OFF/ON | Boolean | Speaker enable signal for PAS cell 1
| Speaker Enable 2 |  | OFF/ON | Boolean | Speaker enable signal for PAS cell 2
| Speaker Enable 3 |  | OFF/ON | Boolean | Speaker enable signal for PAS cell 3
| Speaker Enable 4 |  | OFF/ON | Boolean | Speaker enable signal for PAS cell 4
| Tcell, 0 | degrees C | 0-50 | Numeric | TEC temperature setpoint for PAS cell 0
| Tcell, 1 | degrees C | 0-51 | Numeric | TEC temperature setpoint for PAS cell 1
| Tcell, 2 | degrees C | 0-52 | Numeric | TEC temperature setpoint for PAS cell 2
| Tcell, 3 | degrees C | 0-53 | Numeric | TEC temperature setpoint for PAS cell 3
| Tcell, 4 | degrees C | 0-54 | Numeric | TEC temperature setpoint for PAS cell 4
| Tenable, 0 |  | OFF/ON | Boolean | TEC enable for PAS cell 0
| Tenable, 1 |  | OFF/ON | Boolean | TEC enable for PAS cell 1
| Tenable, 2 |  | OFF/ON | Boolean | TEC enable for PAS cell 2
| Tenable, 3 |  | OFF/ON | Boolean | TEC enable for PAS cell 3
| Tenable, 4 |  | OFF/ON | Boolean | TEC enable for PAS cell 4
Ozone | O2 Valve |  | CLOSED/OPEN | Boolean | O2 valve position
| O3 Valve |  | CLOSED/OPEN | Boolean | Ozone valve position
General | Inlet Valve |  | CABIN/AEROSOL | Boolean | Aerosol path control
| Filter Valve |  | FILTER/SAMPLE | Boolean | Filter path control
| Stop |  |  | Boolean | Non-latching button to control system state
| Save to File |  | NO SAVE/SAVE | Boolean | Latching button to indicate file saving
| Save Ringdowns |  | NO SAVE/SAVE | Boolean | Latching button to indicate ringdown saving
| System messages |  |  | String | System messages produced by the server that indicate current operational state.

## Configuration
Section | Key | Type | Description
Ozone | O3 Power | Boolean | Determines whether the O3 generator is on at boot.
| fO3 | Integer | Number of calibration levels
| ncal_lev | [Integer] | Pressure in mb at each calibration level
| QO3 | [DBL] | Flows to run at each calibration level in lpm
| twait | Integer | Wait time in seconds between each calibration point
| tacq | Integer | Number of seconds to acquire data over for each calibration point
| mode | Boolean | Calibration mode; TRUE = automatic, FALSE = manual
| twarm_u | Integer | Number of seconds required for ozone lamp warm up
MFC | Q0 | [DBL] | Controller setpoints.
Filter | fpos | Boolean | Filter position (TRUE = filter; FALSE = sample
| period | Integer | Number of seconds between the start of speaker cals
| length | Integer | Number of seconds that system is on filter
| auto_cycle? | Boolean | TRUE = system automatically runs filter calibrations
CRDS | flaser | [Integer] | Laser repetition rate in Hz
| Dclaser | [DBL] | fraction of time laser is on in cycle
| Coadd | [Integer] | Number of points to coadd for each cell
| lambda | [DBL] | Laser wavelength in nm for each cell
| RH | [DBL] | RH setpoint for humidified cells
| lcell | [DBL] | Length of cell for each cavity in meters
| expFit | [Str] | Exponential fit method for each cell
| PMTGains | [DBL] | PMT gains for each cell (in Volts)
PAS | Vspeak | [DBL] | Speaker gains for individual cells in Volts
| m | [DBL] | Calibation slope for converting IA to absorption
| b | [DBL] | Calibration offset for converting IA to absorption
| Tcell | [DBL] | Cell temperatures for cells in degree Celsius
| las_pow | [Boolean] | Sets laser power for each cell laser; TRUE = laser on
| tec_pow | [Boolean] | Sets TEC power for each cell; TRUE = TEC on
| chirp.center | Integer | Center frequency in Hz for the chirp
| chirp.band | Integer | Chirp bandwidth in Hz
| spk.cycle | Boolean | Selects whether cycle is automatic; TRUE = automatic
| spk.period | Integer | Seconds between speaker calibration starts
| spk.length | Integer | Length of speaker calibration
TEC | T0 | DBL | Setpoint of TEC in degrees Celsius
| P | DBL | Proportional gain for control
| I | DBL | Integral gain for control
| D | DBL | Derivative gain for control
File | Path | String | Current data path
| Write? | Boolean | Indicates whether file is being written to

## Web Services
Resource | VI | Test Input | Software Response | Actor Responsible | Hardware Response | Associated Task | Testing Device
General | UpdateFilter | 0 | Filter state changed to FALSE | Controller | Voltage on pins 52,18 slot 7 (PXI 6225) goes to 0. | Filter | Multimeter | SCB-68A
|  | 1 | Filter state changed to TRUE | Controller | Voltage on pins 52,18 slot 7 (PXI 6225) goes to 1. | Filter | Multimeter | SCB-68A
| SetTime | 3501638378 | Chassis time set to Dec 17, 2014 05:19:40 | N/A |  |  | Observed in MAX
| NewFile |  | New file generated and previously opened file closed. | File Actor |  |  | Observed in software
| Stop |  | Controller sends stop signal to child actors and system shuts down | Controller |  |  | Observed in software
| FilterCycle | Length=5 Period=10 Autocycle = 1 | Filter switch settings set as defined in command and autocycle is set to TRUE. | Controller | Voltage on pins 52,18 slot 7 (PXI 6225) cycles between 0 and 5 every 5 seconds. | Filter | Multimeter | SCB-68A
CRDS_CMD | fblue | 2000 |  | Instr Actor | 2 kHz pulse generated on ctr0 of slot 8 (PXI 6225). Ctr1 output on same card remains synchronized. | CRDS Blue Laser Output | Oscilliscope | BNC 2110
|  | 3000 |  | Instr Actor | 3 kHz pulse generated on ctr0 of slot 8 (PXI 6225). Ctr1 output on same card remains synchronized. | CRDS Blue Laser Output | Oscilliscope | BNC 2111
|  | 4000 |  | Instr Actor | 4 kHz pulse generated on ctr0 of slot 8 (PXI 6225). Ctr1 output on same card remains synchronized. | CRDS Blue Laser Output | Oscilliscope | BNC 2112
|  | 500 |  | Instr Actor | 500 Hz pulse generated on ctr0 of slot 8 (PXI 6225). Ctr1 output on same card remains synchronized. | CRDS Blue Laser Output | Oscilliscope | BNC 2113
| fred | 500 |  | Instr Actor |  | CRDS Red Laser Output
|  | 2000 |  | Instr Actor
|  | 4000 |  | Instr Actor
| Vpmt | 2,2,2,2,2 |  | Instr Actor | Sets voltages on VCh5,6,7,8,9 to 2 V on PXI-6704. | PMT Gains
|  | 1,2,3,4,5 |  | Instr Actor | Sets voltages on VCh5,6,7,8,9 to 1,2,3,4 and 5 V respectively on PXI-6704. | PMT Gains
| Spk | 100,1300
| UpdateSpkVparams | 0, 1 |  |  | 20 Vpp, 0 centered
|  | 0,0.5 |  |  | 10 Vpp, 0 centered
|  | 0,0.1 |  |  | 2 Vpp
|  | 0,0.05 |  |  | 1 Vpp
|  | 0,0.01 |  |  | 0.2 Vpp
|  | 16384,0.5 |  |  | 10 Vpp, 5V centered
|  | 8192,0.25 |  |  | 10Vpp, 2.5V
|  | 3277,0.1 |  |  | 2Vpp, 1V
|  | 1638, 0.1 |  |  | 2Vpp, 0.5V
|  | 1638,0.05 |  |  | 1Vpp,0.5V
| UpdateLaserEnable | 37 |  |  | All DIO at ~3.3 V
|  | 1 |  |  | Only first at 3.3V

## Sheet3
Resource | VI | Variables | Types | Description | Tested?
General | Data | N/A |  | Returns a JSON formatted data set. | Y
| Stop | N/A |  | Issues a command for system shutdown | Y
| Save | Save | U8 | Toggle save state. Non-zero integers indicate the state should be saving. | Y
| UpdateFilter | State | U8 | Toggle filter state. Non-zero integers set the filter DO high indicating filtering. | Y
| NewFile | N/A |  | Close the current open file and start a new one. | Y
| FilterCycle | Length | U32 | Sets the length of a filter calibration, the period between filter calibrations and the automation state (non-zero is no automation). | Y
|  | Period | U32
|  | Auto | U8
| SetTime | time | DBL | Set the time on the controller to the time specified by the variable. | Y
CRDS_CMD | fblue | Rate | U16 | Set the blue laser repetition rate in Hz. | Y
| fred | Rate | U16 | Set the red laser repetition rate in Hz. | Y
| Vpmt | V | STR | Comma delimited set of voltages for PMTs. Should be one per cell. | Y
PAS_CMD | Spk | df | U16 | Set the characteristics of the speaker chirp for calibration of f0 and Q. df is the bandwidth in Hz and f0 is the center frequency in Hz. | Y
|  | f0 | U16
| SpkSw | SpkSw | U8 | Set the current state of the speaker. A non-zero value will set the speaker DO to high indicating that the speaker is on. | Y
| UpdateFr | f0 | STR | Comma delimited set of center frequencies for setting the modulation frequency. These values will be overridden by speaker calibrations.
| UpdateLaserEnable | LasEnByte | U8 | Single byte that sets the value of the laser enable state for the PAS lasers. 0 will set all lasers off.
| UpdateVrange | Vrange | STR | Comma delimited set of voltages that indicate the range of the modulated laser signal.
| UpdateVoffset | Voffset | STR | Comma delimited set of voltages that indicate the offset of the modulate laser signal.
| UpdateTECsp | TECsp | STR | Comma delimited set of temperatures for setting the cell TEC temperature setpoints on the PAS.
| UpdateTECEn | TECEnByte | U8 | Single byte that sets the value of the PAS cell TEC Enable state. 0 sets all to off.
| UpdateiSpkEn | iSpkEnByte | U8 | Single byute that sets the enable state of the cell speakers. 0 sets all to off. | Y
| UpdateSpkVparams | Voffset | DBL | Sets the voltage parameters for operation of the PAS speakers. | Y
|  | Vrange | DBL
| UpdateSpkCycle | Length | U32 | Set the length of a speaker calibration in seconds, the period between calibrations in seconds and the autocycle option (non-zero values indicate auto). | Y
|  | Period | U32
|  | Cycle | U8

## CVT Entries
PAS | tec.t | [SGL]
| tec.enable | [BOOL]
| spk.cycle | BOOL
| spk.period | I16
| spk.length | I16
| spk.enabled | BOOL
| spk.ienabled | [BOOL]
| las.vrange | [SGL]
| las.voffset | [SGL]
| las.f0 | [SGL]
| las.mod | [U8]
| las.en | [BOOL]
| chirp.fcenter | I16
| chirp.df | I16
| spk.vrange | DBL
| spk.voffset | DBL
CRDS | fblue | I16
| fred | I16
| pmt_gain | [SGL]
General | filter.pos | BOOL
| filter.period | I16
| filter.length | I16
| save | BOOL
| filter.cycle | BOOL
MFCs | Q
TEC | T
P
I
D
Ozone | o2.valve
o3.valve
o3.gen
RHctl | RH
P
I
D

## localStorage
server | ip address of server
port | port for web service access